### SERAPHIN®

### GRADUATED CYLINDER

(WITH NO BOTTOM DRAIN)

### OPERATING AND MAINTENANCE INSTRUCTIONS

### Seraphin® Test Measure

A Division of Pemberton Fabricators, Inc.

Rancocas, NJ

Series F

## JULY, 2004

NOTICE

PEMBERTON FABRICATORS, INC. (SERAPHIN® DIVISION) SHALL NOT BE LIABLE FOR TECHNICAL OR EDITORIAL ERRORS IN THIS MANUAL OR OMISSIONS FROM THIS MANUAL. SERAPHIN®  MAKES NO WARRANTIES, EXPRESSED OR IMPLIED, INCLUDING THE IMPLIED WARRANTIES OF MERCHANTABILITY AND FITNESS FOR A PARTICULAR PURPOSE WITH RESPECT TO THIS MANUAL AND, IN NO EVENT, SHALL SERAPHIN® BE LIABLE FOR ANY SPECIAL OR CONSEQUENTIAL DAMAGES INCLUDING, BUT NOT LIMITED TO, LOSS OF PRODUCTION, LOSS OF PROFITS, ETC.

PRODUCT NAMES USED HEREIN ARE FOR MANUFACTURER OR SUPPLIER IDENTIFICATION ONLY AND MAYBE TRADEMARKS/REGISTERED TRADEMARKS OF THESE COMPANIES.

THE CONTENTS OF THIS PUBLICATION ARE PRESENTED FOR INFORMATIONAL PURPOSES ONLY, AND WHILE EVERY EFFORT HAS BEEN MADE TO ENSURE THEIR ACCURACY, THEY ARE NOT TO BE CONSTRUED AS WARRANTTES OR GUARANTEES, EXPRESSED OR IMPLIED, REGARDING. THE PRODUCTS OR SERVICES DESCRIBED HEREIN OR THEIR USE OR APPLICABILITY. WE RESERVE THE RIGHT TO MODIFY OR IMPROVE THE DESIGNS OR SPECTFTCATIONS OF SUCH PRODUCTS AT ANY TIME.

SERAPHIN® DOES NOT ASSUME RESPONSIBILITY FOR THE SELECTION, USE OR MAINTENANCE OF ANY PRODUCT. RESPONSIBILITY FOR PROPER SELECTION, USE AND MAINTENANCE OF ANY SERAPHIN®  PRODUCT REMAINS SOLELY WITH THE PURCHASER AND END-USER.

SERAPHIN® IS THE REGISTERED TRADEMARK OF PEMBERTON FABRICATORS, INC

BY SERAPHIN® DIVISION of

PEMBERTON FABRICATORS, INC.

RANCOCAS, NJ, U.S.A.

# No part of this work may be reproduced or copied in any form or by any means – graphic, electronic or mechanical- without first receiving the written permission or Pemberton Fabricators, Inc.

CARE, MAINTENANCE, AND USE OF A SERAPHIN® GRADUATED CYLINDER

- A SERAPHIN® graduated cylinder is designed and constructed according to NIST handbook 105-3, OIML R-120 and API Manual of Petroleum Measurement Standards (MPMS) Chapter 4.

- Graduated Cylinder calibration may be established at 1-year intervals and extended or reduced on historical evidence up to the limit determined by your regulatory agency, but should not exceed 3 years due to possible leaking seals which are often difficult to observe under field conditions.

- The graduated cylinder should be maintained in a clean, dent free condition for accurate readings.

I.	Cylinder Wetting Procedure:

- If the graduated cylinder is clean and free of debris, fill the cylinder with liquid from the device being checked.  Fill the cylinder until the amount on the devise reads the nominal capacity of the graduate.

NOTE:  The cylinder must be filled once and drained in order to wet the internal sides of the cylinder body.  Wetting the internals of the cylinder helps ensure consistency in the amount of liquid that clings to the sides, from one test to the next.

- Establishing a “wet-down” condition: Fill the cylinder with product to its nominal level on the scale. Pick up the graduate and inverting it to within 15 degrees of vertical. Once the main flow breaks, drain the cylinder for an additional 10 seconds.

- This constitutes a “wet down” condition and is ready for an official test.

WARNING:  Use of petroleum products require the cylinder be properly

grounded  to prevent possible static electricity sparking which could ignite a

flammable  product. Follow all safety measures.

  - Testing Procedure:

- Record the meter register, or zero the counter if one is being used.

- If a flammable liquid is being tested, ensure the cylinder is grounded          (to eliminate static electricity build-up).

- Ensure that the cylinder is level.

- Fill the cylinder using the dispensing equipment until the meter being checked reads the nominal capacity of the cylinder.

- Identify the volume measurement on the scale by reading the meniscus in the gauge glass. Look directly from eye level to a horizontal line tangential to the bottom of the meniscus. Record the result.

- Pick up the cylinder and inverting it to within 15 degrees of vertical. Once the main flow breaks, drain the cylinder for an additional 10 seconds. This is necessary to provide a consistent “to deliver” condition.

- Perform a second measurement by repeating the procedure in 3. through 6. above.

- If excessive disagreement is observed, check the cylinder for leakage and cleanliness and repeat the procedure until the results of two successive runs repeat within acceptable tolerance.

III.     Care and Maintenance

1. Check cylinder periodically for dents and leaks at the seams and gauge assembly. If a leak is detected, call SERAPHIN® for authorized repair parts.

# Terminology

Bottom loading: Method of filling a volumetric vessel. Intake is made with a bottom load adapter.

Capacity, nominal: The nominal capacity of a field standard test measure or prover is the volume used to designate the measure or prover. The volume is determined by the nominal mark on a graduated upper neck gauge and between the nominal mark on the graduated upper neck gauge and the lower shut off valve or zero mark on a lower neck gauge on a prover.

NOTE:  The nominal capacity of the prover has be set at the Seraphin® facility using Standards that are traceable to the National Institute of Standards and Technology (NIST). If uncertainty values or a higher order of calibration is needed the unit should be sent to a certified calibration laboratory.

Cubical coefficient of thermal expansion: Three dimensional expansion or contraction of a material due to temperature change, expressed °C-1 or °F-1

Field standard test measure:  A measure that can be hand held and is usually less than 40 Liters (10 Gallons).

"High Resolution” standard: A standard with a small diameter neck for improved resolution in reading the meniscus. Generally used in the laboratory as a standard or check standard for measurement control of a primary standard.

Main flow cessation: The moment when a full discharge stream "breaks" and becomes a trickle or a drip.

Prover: Bottom drain is implied. Filled from the top or bottom loading, depending on intended use. May be free standing, mounted permanently, or on truck or trailer and not hand held.

Reference temperature: The temperature at which the measure is intended to contain or to deliver its nominal capacity.

Sight-flow-indicator: A fitting with windows to visually observe the flow through a pipe.

Submerged fill pipe: Pipe used in top filling to minimize foaming of liquids, such as fuel oil and milk, by discharging the product into the bottom of a prover.

To contain:  An indication that the standard is adjusted to contain its intended volume when filled from its empty condition at a reference temperature. (the empty condition is "dry" and test measures or provers are generally not used in this condition).

To deliver:  An indication that the standard is adjusted to deliver its intended volume at a reference temperature. Provers used in a "wet" condition are marked To Deliver.

Tolerance:  Maximum permissible error. A value fixing the limit of allowable error or departure from the true performance or value.

Vapor recovery:  A system for entrapping and collecting vapors for return to the tank to prevent expulsion into the atmosphere.

# SAFETY

The use of a test measure or prover standard may involve hazardous materials, operations, and equipment. Seraphin does not purport to address all safety problems associated with the use of each product. It is the responsibility of the user of the standard to establish appropriate safety and health practices and determine the applicability of local and federal regulatory limitations prior to use. Specific safety information is documented in the various trade references (e.g. American Petroleum Institute and Petroleum Equipment Institute).

Commercial liquid measuring devices, tested with provers, are typically used to measure quantities of petroleum products. Petroleum products are known hazardous materials and hazardous wastes. The user is encouraged to obtain Material Safety Data Sheets (MSDS) from the manufacturer of any product encountered. Federal, state and local safety and disposal regulations concerning hazardous materials encountered should be reviewed by the user.

Safety devices and locks should be installed to prevent inadvertent operation of, or unauthorized tampering with, equipment. All automated or power-operated meter proving systems should have emergency manual operators for use during an accident or power failure. Grounding devices should be provided to protect against electrical shock or static discharge in both tank prover and \ electrical instrumentation.

All electrical connections must be explosion proof. All wiring, including low voltage wiring shall meet the requirements of Article 300, 500, Group D, Class 1, Division 1, and  250.45 and/or other applicable articles of the latest edition of the National Electrical Code.

SERAPHIN® TEST MEASURES

DIVISION OF PEMBERTON FABRICATORS, INC.

PO BOX 227     30 INDEL AVENUE

RANCOCAS, NJ  08073-0227

# CERTIFICATE  OF  WARRANTY

PEMBERTON FABRICATORS, INC. FULLY WARRANTS THAT THE PRODUCT SUPPLIED SHALL CONFORM TO THE DESCRIPTION IN THE QUOTATION AND AGREES TO REPAIR OR REPLACE F.O.B. SHIPPING POINTS, ANY PARTS EXCEPTING EXPENDABLE ITEMS SUCH AS FILTERS AND LIGHTS THAT FAIL DUE TO DEFECTS IN MATERIAL OR WORKMANSHIP WITHIN ONE YEAR OF USE OR 18 MONTHS AFTER SHIPMENT, WHICHEVER OCCURS FIRST.

Other than those expressively stated herein, there are no other warranties of any kind, expressed or implied, and specifically excluded but not by way of limitation, are the implied warranties of fitness for particular purpose and merchantability.

It is understood and agreed that Pemberton Fabricators, Inc.’s liability whether in contract, in tort, under any warranty, in negligence or otherwise shall not exceed the cost of repair or replacement, f.o.b. shipping points, of defective parts.  Under no circumstances shall Pemberton Fabricators, Inc. be liable for special, indirect, incidental, or consequential damages.  The price stated for the equipment is a consideration in limiting Pemberton Fabricators, Inc.’s liability.  No action, regardless of form, arising out of the transactions of this agreement may be brought by purchaser more than one year after the cause of action has accrued.

The warranty for the equipment proposed is as stated in the above paragraphs.  It is not re-stated--nor does it appear in any other form.